Analog output
Power supply range : 10V~30V DC (0-10V output supply voltage needs 24VDC)
Output : 4-20mA, 0-5V, 0-10V
Maximum power consumption
Current output: 0.08W
Voltage output: 0.08W
Working humidity: 0%~100%RH
Working temperature :-40 C ~ 60 C
Spectral range :0.3~3μm
Measuring range: 0~1800W/m2
Resolution :1W/m2
Response time <=10S
Nonlinear <+-2%
Annual stability <+-2%
Cosine response <+-10%
Line length 60cm

Installation method
1. After the installation is complete, remove the protective cover
2. Make sure the bracket is installed and the radiation sensor is parallel to the ground.
3. Secure the sensor to the mounting bracket by using a screw through the mounting hole on the sensor.
4. Please note that the dust cover should not be damaged during the installation process, so as not to affect the measurement accuracy.

Wiring
	Wire color	Description
Power	Brown	Power supply (10~30V DC)
	black	Power-
Communication	blue	Signal+
	green	Signal-

Calculation formula
3.1 Current type signal output calculation method
For example, the range is 0~1800W/m2, 4~20mA output. When the output signal is 12mA, the current solar radiation value is calculated.
The solar radiation range spans 1800 W/m2 and is expressed by a 16 mA current signal, 1800 W/m2/16 mA = 112.5 W/m2/mA, that is, current 1mA represents solar radiation change 112.5W/m2, measured value 12mA-4mA=8mA, 8mA*112.5W/M2/mA = 900W/m2, the current solar radiation value is 900W/m2.
3.2 Voltage type signal output conversion calculation
For example, the range is 0~1800W/m2, 0-10V output. When the output signal is 5V, the current solar radiation value is calculated. sun
The span of the radiation range is 1800W/m2, expressed by a 10V voltage signal, 1800W/m2/10V=180W/m2/V, that is, electricity
Pressure 1V represents solar radiation change 180W/m2, measured value 5V-0V=5V, 5V*180W/m2/V=900W/m2, current sun
The radiation value is 900 W/m2.

RS485 Output

Power supply range :10V~30V DC
Output : RS 485 (standard Modbus-RTU protocol)
Power consumption :0.08W
Working humidity: 0%~100%RH
Working temperature: -40 C ~ 60 C
Spectral range: 0.3~3μm
Measuring range : 0~1800W/m2
Resolution :1W/m2
Response time :≤ 10S
Nonlinear: <±2%
Annual stability :≤ ± 2%
Cosine response: ≤ ± 10%
Line length : 60cm

Installation method
1. After the installation is complete, remove the protective cover
2. Make sure the bracket is installed and the radiation sensor is parallel to the ground.
3. Secure the sensor to the mounting bracket by using a screw through the mounting hole on the sensor.
4. Please note that the dust cover should not be damaged during the installation process, so as not to affect the measurement accuracy.

Wiring
	Wire color	Description
Power	Brown	Power supply (10~30V DC)
	black	Power-
Communication	yellow	485+
	blue	485-

how to change device address and baud rate ?
1. pls choose the right COM port . ( my computer-Device Manager-port(COM and IPT )
check
2. connect a PC and click software to check the baud rate ( 4800bit/s ) Address ( 0X01 )
3. change thge baud rate and the address you need .

1. communication protocol
Code	8 bit binary system
Data bits	8bits
parity check bit	NO
stop bit	1BIT
Error checking	CRC
baud rate	2400bit/s、4800bit/s、9600 bit/s （4800bits by default）

2. Using Modbus-RTU communication protocol, the format is as follows:
The initial structure is more than 4 bytes of time
Address code = 1 bytes
Function = 1 bytes
Data area = N bytes
Error check = 16 bit CRC code
The end of the structure is greater than or equal to 4 bytes of time
Address code: the address of the transmitter, which is unique in the communication network (factory default 0x01).
Function code: the instruction function instruction from the host. The transmitter only uses the function code 0x03 (read register data).
Data area: data area is specific communication data, pay attention to 16bits data before the high byte!

Host query frame structure:
Address code	Function code	Register start address	Register length	check code low	Check code high
1 bytes	1 bytes	2 bytes	2 bytes	1 bytes	1 bytes

Slave transponder frame structure
Address code	Function code	effective byte	No 1 Data area	NO2 Data area	No N Data area	Check code
1 bytes	1 bytes	1 bytes	2 bytes	2 bytes	2 bytes	2 bytes

Register address
Register address	content	Operation (hexadecimal)	Specification and definition
0000 H	Solar radiation value	03	actual value
0052H	Deviation	03/06	Solar radiation deviation / (0~1800)
07D0 H	Device address	03/06	1~254 (factory default 1)
07D1H	Device baud rate	03/06	0 represents 2400 / 1 represents 4800 / 2 represents 9600

Communication protocol example and explanation
4.4.1 Reading the current solar radiation value
Inquiry frame

Address code	Function code	Start address	Data length	check code low	Check code high
0x01	0x03	0x00 0x00	0x00 0x01	0x84	0x0A

Response frames
Address code	Function code	effective byte	Solar radiation value	check code low	Check code high
0x01	0x03	0x02	0x00 0x64	0x9B	0xAF

Solar radiation value:
0064 (hexadecimal) =100=> solar radiation value = 100W/m2

Write bias value
Inquiry frame

Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x00 0x00	0x00 0x0A	0x09	0xCD

Response frame
Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x00 0x00	0x00 0x0A	0x09	0xCD

Write current solar radiation deviation value
000A (hexadecimal) =10=> Solar radiation deviation value = 10W/m2 The deviation value is 10W/m2

Modify current address
Inquire frame (modify current address is 0x02)
Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x07 0xD0	0x00 0x02	0x08	0x86

Response frame
Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x07 0xD0	0x00 0x02	0x08	0x86

Modify the current baud rate
Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x07 0xD1	0x00 0x02	0x59	0x46

Response frame
Address code	Function code	Register start address	Modify value	check code low	Check code high
0x01	0x06	0x07 0xD1	0x00 0x02	0x59	0x46

Query current address
Inquiry frame
Address code	Function code	Start address	Data length	check code low	Check code high
0xFF	0x03	0x07 0xD0	0x00 0x01	0x91	0x59

Response frame
Address code	Function code	Returns the number of valid bytes	address	check code low	Check code high
0xFF	0x03	0x02	0x00 0x01	0x50	0x50

The address code read is the real address of the device: 01

Precautions and troubleshooting
Precautions:
1. When the customer receives the product, please confirm the product model, etc.
2. Do not electrify the wiring. After the wiring check is correct, the power can be turned on.
3. The sensor is a precision device. Do not disassemble the protective transparent cover.
Troubleshooting:
1. If the read value is displayed as 0, check if the product protection cover is removed.
2. Please check if the 485 wiring is correct.
3. Check if the power supply meets the label
4. Equipment damage
6. Product maintenance
1. The dust cover should be kept clean, often with a soft cloth or wipe
2. There should be no water in the dust cover. If you encounter heavy rain, snow, ice, etc. for a long time, it is recommended to cover it.